TOTAL BILIRUBIN TEST

(TECHNICAL PROCEDURE)

Bilirubin Test

A bilirubin test measures the levels of bilirubin in your blood. Bilirubin is a yellowish pigment made during the breakdown of red blood cells, which passes through the liver and is eventually excreted out of the body.

Principle

A total bilirubin test is usually one of the tests that check the health of your liver. It is done to:

- Investigate jaundice — a yellowing of the skin and eyes caused by high bilirubin levels.

- Determine whether there might be a blockage in the bile ducts in the liver or the gallbladder.

- Help detect liver disease, particularly hepatitis, or monitor its progression.

- Help evaluate anemia caused by the destruction of red blood cells.

- Help follow how a treatment is working.

- Help evaluate a suspected drug toxicity.

In this procedure, total bilirubin, in the presence of a solubilizing agent, is coupled with 3,5- dichlorophenyl diazonium in a strongly acidic medium. The intensity of the red azo dye formed is directly proportional to the total bilirubin and can be determined photometrically (546 nm).

Specimen requirements

- Specimen Type: Serum, lithium heparin plasma, and K2-EDTA plasma. Other anticoagulants and grossly hemolyzed blood samples are not acceptable. The specimen should be protected from light.

- Labeling: Patient name, patient medical record number with check digit, collection date and time, specimen type, test required

- Collection Tube:

- Preferred: Serum Separator Tube (SST)

- Alternative: Red Top Blood Tube

- Submission Container/Tube:

- Amber vial (used to protect the sample  from UV light)

- Collection instructions:

1. ) Serum gel tubes should be centrifuged within 2 hours of collection.

2.) Red-top tubes should be centrifuged, and the serum should be aliquoted into an amber vial within 2 hours of collection.

- Specimen Stability:

- Specimen Type - Serum

  - Refrigerated (preferred) - 7 days, Light Protected Container

  - Frozen - 30 days, Light Protected Container

  - Ambient - 24 hours, Light Protected Container

Reagents or Media, Supplies, Equipment

- Reagents and Supplies: Roche Cat. No. 05795397190, BILT3 reagent kit (250 tests)

- Reagent composition (Ready to use; no preparation required):

  - R1 reagent: Phosphate: 25 mmol/L; detergent; stabilizers; pH1.0

  - R2 reagent: 3,5-dichlorophenyl diazonium salt: >1.35 mmol/L

Storage and stability. Keep reagents stored at 2-8°C until use. The reagents are stable for 12 weeks refrigerated on the analyzer.

- Equipment/Instrumentation

  - Roche Cobas 6000 Chemistry Analyzer

Calibration

- The calibrator is stable until the bottle expires when stored at 4ºC.

  - The lyophilized calibrator is prepared with 3.0 mL of deionized water.

  - Pipette the water into the bottle, and then dissolve by gentle swirling within 30 minutes.

  - Avoid the formation of foam while mixing.

  - The prepared calibrator is stable for eight hours at room temperature, two days at 4ºC, and one month at –20ºC (frozen once).

- Calibration frequency: When a reagent lot number changes, a two-point calibration (H2O + C.F.A.S.) must be performed. The Cobas 6000 will not allow testing to proceed until a successful calibration has been completed. Monitor control values to determine the stability of the current calibration.

- Manual calibration should be performed if:

  - A reagent lot change has not occurred in the past 6 months

  - After significant service or repairs

  - As needed for troubleshooting

- If calibration fails, perform the following corrective action steps in sequence:

  - Check the reagent and calibrator for appropriate lot numbers, expiration dates, and preparation and storage conditions.

  - Repeat calibration with a new calibrator.

  - Repeat calibration with new reagent and new calibrator

Quality Control

- Normal Pooled Serum Control (CQ)

  - Storage Stability:

  - It can be stored at -80°C for up to 4 years.

  - Remains stable at refrigerated temperatures (2-8°C) for up to 1 day.

  - It can be kept at room temperature for up to 4 hours.

- Roche Precipath U Plus Control (Quality control material)

  - Storage Stability: Stable until the expiration date on the package when unopened and stored at 2-8°C.

  - Preparation:

- Use two bottles.

- Open Bottle 1.

- Pipette precisely 3.0 mL of diluent from Bottle 2 into Bottle 1.

- Gently swirl for 30 minutes to dissolve.

  - Stability of Prepared Control:

- 12 hours at room temperature.

- 5 days at 2-8°C.

- 1 month at -20°C (after being frozen once).

- Quality Control Procedures

  - Testing Frequency:

- Quality control samples are analyzed at the start of each day and at the end of each shift.

- Additional checks are done after changing reagents, following major maintenance, or whenever troubleshooting is needed.

- Analytical Measurement Range (AMR) Validation

  - The AMR must be validated every 6 months or after significant maintenance or service.

- New Lot Verification

  - Each new reagent lot must be verified for acceptability before use, which includes calibration, quality control, and comparing at least 5 patient samples between the old and new lots.

Step-by-step instructions

- Quantitative Testing

- Principle: Total bilirubin in serum or plasma is measured by coupling with 3,5-dichlorophenyl diazonium in an acidic medium, forming a red azo dye that is photometrically analyzed at 546 nm.

- Instrument: Roche Cobas 6000, the chemistry analyzer, analyzes the samples.

- Specimen Requirements: Serum, lithium heparin plasma, and K2-EDTA plasma (100 µL optimum volume). Specimens should be protected from light to prevent bilirubin degradation.

- Qualitative Testing:

  - The method used does not specify qualitative testing for bilirubin, as it is a quantitative procedure.

- Interpretation of Results:

  - Reportable Range: 0.1–175.5 mg/dL.

  - Results are reported to one decimal place and interpreted based on clinical relevance, with critical values determined by disease-specific factors.

Reporting Results

- Reference Intervals:

  - Serum, adult: Normal values are up to 1.0 mg/dL.

- Procedures for Reporting Abnormal Results:

  - If the test values exceed the pre-programmed thresholds, automatic re-analysis is triggered. If results remain abnormal after repeat testing, they are reported as abnormal.

- Reporting Format:

  - Results are reported electronically in mg/dL, with low results as <0.1 mg/dL. Errors or flagged values must be checked and corrected as needed.

Procedure Notes

- Special Precautions:

  - Handle all specimens as though capable of transmitting infectious diseases.

  - Protect bilirubin samples from light to avoid degradation.

- Possible Sources of Error:

  - Grossly hemolyzed samples can interfere with readings.

  - If samples are not protected from light.

- If calibration fails, check for reagent preparation, stability, and potential system issues. If necessary, contact technical support.

Limitations of Methods

- Detection Limit: 0.1 mg/dL.

- Interferences: A hemolytic index below 800 and a lipemic index below 1000 show no interference. Certain drugs, like Hydroxocobalamin, may cause false readings.

A troubleshooting or backup plan

- Stop testing immediately when the system fails and notify the supervisor.

- Check calibration -  Verify the reagent lot numbers, expiration dates, and preparation/storage conditions to ensure they are correct.

- Recalibrate the system using fresh calibrators and reagents. Repeat the analysis if needed.

- If calibration or system problems persist,

  - Contact the manufacturer of the equipment

  - The usage of spectrophotometers might serve as a temporary backup if manual testing is needed.

  - Deliver the specimens to another laboratory, a reference laboratory, or a partner laboratory,  equipped to run the test

References

- Bilirubin test - Mayo Clinic. (n.d.). https://www.mayoclinic.org/tests-procedures/bilirubin/about/pac-20393041

- Bilirubin Total and Direct - Laboratory Test Directory | West. (n.d.). https://corewellhealth.testcatalog.org/show/LAB168-1

- Specimen Labeling Requirements | Department of Pathology | School of Medicine | University of California, Irvine. (2023, May 17). https://www.pathology.uci.edu/services/specimen-labeling-requirements.asp

- Total and Direct Bilirubin. (n.d.). Biolabo. https://www.biolabo.fr/pdfs/noticesE/biochimieE/AT-80403.pdf

- University of Minnesota. (2018). Serum Total bilirubin Laboratory Procedure Manual. Advanced Research and Diagnostic Laboratory (ARDL). https://wwwn.cdc.gov/nchs/data/nhanes/2017-2018/labmethods/BIOPRO-J-MET-Total-Bilirubin-508.pdf

Effective Date

October 8, 2024

Laboratory Director

ATHEA FRANCHESCA S. BAYNO